## Slide 1
Team GMora
BRAINSTORM 2026 · SEMIFINAL REPORT
PneumoScan AI
Explainable, fail-safe pneumonia screening from chest X-rays — built, validated, and deployed for Sri Lankan primary care.
Live at pneu.gmora.dev
94.2% accuracy · 0.988 AUC
Installable mobile app
University of Moratuwa
Kusal Pabasara · Pasindu Mihiranga
Thanoj Buddhima · Senuja Dilmith
Innovative & Affordable Healthcare
Solutions for Local Challenges

## Slide 2
PneumoScan AI
EXECUTIVE SUMMARY
AT A GLANCE
From concept to a deployed clinical screening tool
Since the initial round, PneumoScan AI has progressed from a proof-of-concept model to a fully working, publicly deployed application — with a substantially stronger, independently validated model and a complete patient-and-clinician workflow.
94.2%
TEST ACCURACY
▲ from 86.7%
0.988
AUC-ROC
▲ from 0.943
94.3%
SENSITIVITY
fail-safe tuned
23,405
TRAINING IMAGES
peds + adult
What has been achieved
  A live, installable app at pneu.gmora.dev — served over HTTPS, installable to a phone home screen (PWA), with in-browser camera capture and file upload.
  A validated, generalised model — retrained on a combined pediatric + adult corpus (23,405 images) with dedicated phone-photo robustness, verified by 5-fold cross-validation.
  A fail-safe clinical design — a sensitivity-first decision threshold, an image-quality gatekeeper, calibrated "uncertain" handling, and Grad-CAM explainability on every prediction.
  Two audiences, one engine — tuned plain-language guidance for patients/parents and a full technical read-out for clinicians.
Positive result, verified: On a real phone photo of an adult pneumonia X-ray from a Moratuwa clinic, the system correctly returns "Signs of pneumonia"; on a healthy adult phone X-ray it returns "No clear signs" — the two cases the earlier pediatric-only model could not handle.
How this report is organised
Section 1 frames the problem and local need. Section 2 describes the deployed solution and how it works. Section 3 presents the technical implementation, model, and validated results. Section 4 reviews the literature and background research. Sections 5–8 address feasibility, regulatory compliance and IP, data security, and ethics. Sections 9–10 cover impact, sustainability, cost-effectiveness, and the roadmap.
Team GMora · University of Moratuwa
Brainstorm 2026 · Semifinal

## Slide 3
PneumoScan AI
SECTION 1 · THE PROBLEM
SECTION 1
1. The problem and why it matters
Pneumonia is a leading cause of mortality worldwide, and in Sri Lanka timely diagnosis is blocked by a critical shortage of radiologists — especially in rural and district facilities.
GLOBAL & LOCAL BURDEN
  ~450 million pneumonia cases occur globally each year (WHO).
  ~2.5 million deaths annually — the leading infectious killer of children under five.
  Pneumonia accounts for roughly 15% of all deaths in children under five.
  In Sri Lanka, respiratory infections are among the top causes of hospital admission; rural centres face diagnostic delays of 24–72 hours awaiting radiologist review.
  Radiology is rarely available around the clock, leaving nights and weekends inadequately covered.
UNMET CLINICAL NEEDS
Radiologist shortage
Chest X-rays go unread or delayed in rural Sri Lanka where trained radiologists are scarce.
Diagnostic variability
Manual interpretation carries inter-observer variability, particularly for subtle or early pneumonia.
No affordable pre-screening
No low-cost, AI-assisted triage tool is widely deployed in Sri Lankan hospitals.
24/7 unavailability
Specialist services are rarely continuous, leaving critical gaps in emergency care.
This directly aligns with the Brainstorm 2026 theme — "Innovative and Affordable Healthcare Solutions for Local Challenges." An affordable, AI-assisted screen deployed in rural clinics and district hospitals can shorten dangerous diagnostic delays and prevent avoidable deaths.
HOW WE RESEARCHED THE PROBLEM
  WHO & global health data to quantify the burden of pneumonia in South Asia.
  Peer-reviewed clinical dataset (Kermany et al., 2018; Cell) validating the authenticity of the pediatric training data.
  Expert feedback from the IEEE EMBS BioFusion Hackathon 2026, where domain experts reviewed the project and guided refinement of the model and app design.
Team GMora · University of Moratuwa
Section 1 · The Problem

## Slide 4
PneumoScan AI
SECTION 2 · THE SOLUTION
SECTION 2
2. The solution — a deployed screening app
PneumoScan AI is a live web and mobile application. A health worker — or a patient — captures or uploads a chest X-ray and receives an instant, explainable, fail-safe screening result.
Now live: https://pneu.gmora.dev — HTTPS-secured, installable to a phone home screen, no app store required.
How the app works — end to end
1
Capture or upload. Take a photo of the X-ray with a phone camera, or upload a digital file (JPEG/PNG). A role selector tailors the experience for a patient/parent or a clinician.
2
Quality gate. An automatic check rejects unusable images (too small, no detail) and — for phone photos — tolerates the soft focus and warm colour cast that are normal for a snapshot of a film.
3
Phone-photo rectification. For photos of physical films, OpenCV detects the film boundary and corrects the camera-angle perspective.
4
AI screening. A fine-tuned ResNet-50 outputs a calibrated pneumonia probability; a fail-safe decision threshold converts it into a triage band.
5
Explainable result. A Grad-CAM heatmap shows which lung regions drove the prediction, alongside plain-language guidance and an explicit "screening aid, not a diagnosis" disclaimer.
Result communication — safe by design
Band
Meaning
Guidance shown
No clear signs
Low pneumonia probability
Reassurance + "see a doctor if symptoms persist"
Uncertain
Borderline / on the fence
"Get a second opinion" — never a confident call
Signs of pneumonia
Above fail-safe threshold
Prompt to seek clinical care promptly
For patients & parents
Plain-language result, clear next-step guidance, and a strong prompt to consult a clinician. No jargon.
For clinicians
Grad-CAM overlay, class probabilities, the fail-safe operating point, and full technical detail for decision support.
Team GMora · University of Moratuwa
Section 2 · The Solution

## Slide 5
PneumoScan AI
SECTION 3 · TECHNICAL IMPLEMENTATION
SECTION 3
3. Technical implementation & results
The model was substantially strengthened: broader training data, deeper fine-tuning, phone-photo robustness, and a fail-safe operating point — all validated on held-out data.
Model & training strategy
  Backbone: ResNet-50 (ImageNet-pretrained). We now fine-tune the deepest residual stage (layer4) plus the classifier head — ~15M trainable parameters — rather than the head alone, letting high-level features adapt from natural images to radiographs.
  Combined corpus: pediatric (Kermany/Guangzhou, 5,856) + adult (COVID-19 Radiography Database, 17,549) = 23,405 images, near class-balanced (≈50/50).
  Phone-photo augmentation: training images are synthetically transformed to mimic phone snapshots of films — perspective, warm colour cast, glare, vignette, blur, JPEG — so the model reads real phone captures, not just clean digital scans.
  Regularisation: AdamW with weight decay, label smoothing, cosine LR schedule, early stopping — the train/validation gap stays ≈1%.
  Fail-safe threshold: the decision threshold is tuned on a validation set for ≥95% sensitivity and capped so the tool errs toward flagging — a missed pneumonia is the costly error.
Held-out test performance (3,510 images)
94.2%
ACCURACY
94.3%
SENSITIVITY
94.2%
SPECIFICITY
0.988
AUC-ROC
Metric
Old (pediatric-only)
Now (combined + phone-trained)
Test accuracy
86.7%
94.2%
AUC-ROC
0.943
0.988
Sensitivity / Precision / F1
96.7% / 84.3% / 0.90
94.3% / 94.1% / 0.94
Works on adult X-rays
No (fails)
Yes — 92.4% recall, AUC 0.984
Works on phone photos
No
Yes — verified on real captures
Overfitting (dev→test gap)
7.7%
1.1%
Validation depth. Beyond a single split, we ran 5-fold stratified cross-validation (sensitivity 98.7% ± 0.5% across folds) and report per-population results — pediatric (97.6% recall, AUC 0.995) and adult (92.4% recall, AUC 0.984) — demonstrating the model generalises across ages rather than overfitting one population.
Team GMora · University of Moratuwa
Section 3 · Technical Implementation

## Slide 6
PneumoScan AI
SECTION 3 · ARCHITECTURE
SECTION 3 (CONT.)
System architecture & live deployment
A single trained model serves both audiences through a lightweight, low-cost, fully deployed web stack — evidence of functionality, not just simulation.
Inference pipeline
Capture / Upload → Quality gate (OpenCV) → Phone-photo rectification (perspective correction) → Preprocess (224×224, RGB, ImageNet norm) → ResNet-50 (fine-tuned) → Softmax P(Normal)/P(Pneumonia) → Fail-safe triage band → Grad-CAM heatmap → Role-aware result
Deployment stack (all live)
Layer
Technology
Role
Model
PyTorch 2.6, torchvision ResNet-50
Trained on an NVIDIA RTX 3050 GPU; runs CPU-only inference in production.
Vision utils
OpenCV, NumPy, Pillow
Quality checks, phone-photo rectification, Grad-CAM overlay.
App / UI
Streamlit + custom CSS/PWA
Patient/clinician UI, camera capture, installable home-screen app.
Server
AWS EC2 (Ubuntu), nginx, systemd
Reverse proxy, process supervision, always-on service.
Security
Let's Encrypt TLS (HTTPS)
Encrypted transport; required for browser camera access.
Delivered application features
Live at pneu.gmora.devInstallable PWA In-browser camera capturePhone-photo rectification Image-quality gateFail-safe threshold Calibrated "uncertain" bandGrad-CAM explainability Patient / Clinician rolesClinical escalation guidance Offline app shellModel insights dashboard
Feasibility demonstrated. The entire system runs on free/low-cost infrastructure: open-source frameworks, a single small cloud VM for CPU inference, and no per-query licensing. An optional edge deployment (Raspberry Pi / Jetson Nano) enables fully offline use in clinics without internet.
Team GMora · University of Moratuwa
Section 3 · Architecture & Deployment

## Slide 7
PneumoScan AI
SECTIONS 4–5 · BACKGROUND & FEASIBILITY
SECTION 4
4. Literature review & background
Work
Contribution
Our advance
Rajpurkar et al. (2017) — CheXNet
121-layer DenseNet, radiologist-level pneumonia detection.
We add explainability + a deployed, fail-safe workflow for low-resource use.
Kermany et al. (2018)
Transfer learning on pediatric CXR; expert-graded dataset.
We generalise beyond pediatrics to adults and phone photos.
Selvaraju et al. (2017) — Grad-CAM
Gradient-based visual explanation for CNNs.
Integrated into every prediction to build clinical trust.
He et al. (2016) — ResNet
Residual learning enabling deep networks.
Our fine-tuned backbone, adapted to CXR features.
Phillips et al. — CheXphoto
Photographed-CXR robustness benchmark.
We reproduce capture artefacts synthetically to train for phone use.
Gap addressed: most published models are accurate on clean, in-distribution scans but are not deployment-ready, not explainable, and not validated on the messy inputs (adult anatomy, phone photos) that a real rural workflow produces. PneumoScan AI targets exactly that gap.
SECTION 5
5. Feasibility
Technical feasibility
Built entirely on mature, open-source frameworks (PyTorch, OpenCV, Streamlit). The trained model is ~90 MB and runs in well under a second on CPU — no GPU needed in production.
Operational feasibility
Deployed on a single low-cost cloud VM with automated TLS and process supervision; an edge option removes the internet dependency for remote clinics.
Data feasibility
Trained on 23,405 openly licensed, peer-reviewed images; the pipeline can ingest new regional data as partner hospitals contribute.
Economic feasibility
Zero per-query licensing; runs on free-tier / low-cost infrastructure, making per-scan cost negligible versus enterprise tools.
Team GMora · University of Moratuwa
Sections 4–5 · Background & Feasibility

## Slide 8
PneumoScan AI
SECTIONS 6–7 · REGULATORY & SECURITY
SECTION 6
6. Regulatory aspects & intellectual property
As a healthcare tool, PneumoScan AI is positioned as Software as a Medical Device (SaMD) and framed explicitly as a screening decision-support aid, not a diagnostic device — keeping a clinician in the loop and lowering the regulatory risk class.
Jurisdiction / framework
Applicability & our plan
NMRA (Sri Lanka)
Primary regulator. We target registration as a low–moderate risk SaMD, with a documented clinical-validation study on local X-rays before any clinical deployment.
US FDA
Reference framework: an equivalent product would pursue a 510(k) pathway as Class II. We follow FDA "Good Machine Learning Practice" and SaMD guidance for design controls.
Standards
Aligning with IEC 62304 (medical device software lifecycle) and ISO 14971 (risk management) as the engineering discipline for a future submission.
INTELLECTUAL PROPERTY
  The trained model, phone-photo augmentation pipeline, fail-safe threshold logic, and the patient/clinician application are our original contributions.
  All third-party components (ResNet-50 weights, datasets, libraries) are used under permissive open-source / open-access licences, documented for compliance.
  Strategy: retain proprietary rights over the application and locally-trained weights while keeping the core screening accessible to public hospitals (see sustainability model).
SECTION 7
7. Data security & privacy policy
In transit
All traffic is encrypted over HTTPS (TLS). Camera access is only granted on a secure origin.
Data minimisation
The current tool performs inference and returns a result without persisting patient images by default — nothing is stored unless explicitly enabled with consent.
Consent & retention
Planned account tier gates any storage behind explicit consent, with retention limits and one-tap deletion.
Governance
An audit trail of predictions (not identities) supports quality monitoring and model-drift detection without exposing personal data.
Responsible-AI stance: every result is labelled a screening aid requiring clinician confirmation; the tool is scoped and its known limitations (see ethics) are stated in-app.
Team GMora · University of Moratuwa
Sections 6–7 · Regulatory & Security

## Slide 9
PneumoScan AI
SECTIONS 8–9 · ETHICS & IMPACT
SECTION 8
8. Ethical review
  Safety-first framing. A fail-safe threshold minimises missed pneumonia; a false positive (healthy patient referred for review) is the safe direction and is handled by the "confirm with a doctor" flow.
  Honesty about limits. The model is strongest on clear digital X-rays; performance on phone photos of adult films is stated openly, and borderline cases are shown as "uncertain" rather than a confident call.
  Bias & fairness. Training now spans pediatric and adult populations; per-population metrics are reported. Continued external validation on Sri Lankan data is part of the roadmap.
  Human-in-the-loop. The system never replaces a clinician — it triages and explains, leaving the diagnosis to a qualified professional.
  Accessibility. Bilingual (Sinhala/English) interfaces are planned so the tool serves the communities that need it most.
SECTION 9
9. Impact on the healthcare sector
PneumoScan AI turns a 24–72 hour diagnostic delay into a sub-30-second, always-available assessment, directly addressing radiologist shortage, delay, and the absence of automated pre-screening.
Stakeholder
Benefit
Rural health workers
Rapid AI-assisted pre-screening on a basic smartphone — faster decisions without specialist access.
Patients
Shorter waits, fewer missed cases, earlier treatment — improved outcomes.
District hospitals
A 24/7 second opinion that helps prioritise urgent cases and reduce radiologist load.
Ministry of Health
Scalable, low-cost deployment and a basis for national pneumonia-trend tracking.
Team GMora · University of Moratuwa
Sections 8–9 · Ethics & Impact

## Slide 10
PneumoScan AI
SECTION 10 · SUSTAINABILITY & ROADMAP
SECTION 10
10. Sustainability, cost & roadmap
Cost-effectiveness
Open-source stack, CPU-only inference on a single low-cost VM, and zero per-query licensing make per-scan cost negligible — versus enterprise tools (Qure.ai, Lunit, Aidoc) priced out of reach for rural clinics.
Sustainability model
Freemium access for NGOs and government hospitals; an optional subscription tier for private hospitals funds maintenance. The model improves as partner hospitals contribute regional data.
Roadmap
Phase
Milestone
1 · Done
Deployed binary screening with Grad-CAM, fail-safe threshold, adult + pediatric + phone-photo robustness.
2
Bacterial vs. viral pneumonia classification for targeted treatment guidance.
3
Additional local pathologies (tuberculosis, pleural effusion).
4
SaMD registration with NMRA Sri Lanka + local clinical-validation study.
5
Bilingual UI and integration with national health IT (eSuwa).
Social impact. By making pneumonia screening fast, explainable, and affordable across rural Sri Lanka, PneumoScan AI can help prevent avoidable deaths — bringing specialist-level decision support to the settings that need it most.
Key references
  Rajpurkar P. et al. (2017). CheXNet: Radiologist-Level Pneumonia Detection on Chest X-Rays. arXiv:1711.05225.
  Kermany D. et al. (2018). Identifying Medical Diagnoses by Image-Based Deep Learning. Cell 172(5).
  He K. et al. (2016). Deep Residual Learning for Image Recognition. CVPR.
  Selvaraju R. et al. (2017). Grad-CAM: Visual Explanations from Deep Networks. ICCV.
  Rahman T. et al. (2021). COVID-19 Radiography Database. Kaggle / Computers in Biology and Medicine.
  World Health Organization — Pneumonia fact sheets and global health statistics.
Team GMora · University of Moratuwa
Brainstorm 2026 · Semifinal Report